|  |  |
| --- | --- |
|  | De Mayo Knee Positioners**STERILIZATION PROCEDURE** |

# Recommended Sterilization Instructions

* Insure that all parts are thoroughly cleaned.
* Make sure that all moveable parts are loose and can move freely.
* DO NOT PLACE POSITIONER IN A MILK BATH OR LUBRICATE.
* If using a sterilization case follow Instructions For Use for the product.
* Double wrap in two disposable wraps. Use 48 inch x 48 inch wraps.

(approximately 122 cm x 122 cm).

* Run normal vacuum cycle for your institution
* **STEAM STERILIZATION ONLY- ALL OTHER STERILIZATION METHODS NOT VALIDATED**

**MINIMUM PARAMETERS PRE-VAC STERILIZATION**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Product****Part Number** | **With Sterilization Case****Case Part Number** | **Temperature****Setting** | **Exposure****Time** | **Dry****Time** |
| De Mayo Knee Positioner803 & 803-S | 706-S | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
| De Mayo D2 Knee Positioner803-114 | 706-S | 270°F to 272°F132°C to 134° C | 4 minutes | 30 minutes |
| De Mayo V2 Knee Positioner114 & 114-S | 706-S | 270°F to 272°F132°C to 134° C | 4 minutes | 30 minutes |
| De Mayo V2 E Extension Arm114 - EXT | 706-S | 270°F to 272°F132°C to 134° C | 4 minutes | 40 minutes |
|  | **Without Sterilization Case** |  |  |  |
| De Mayo Knee Positioner803 & 803-S | Without Sterilization Case Double Wrap Only | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
| De Mayo D2 Knee Positioner803-114 & 803-114-S | Without Sterilization Case Double Wrap Only | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
| De Mayo V2 Knee Positioner114 & 114-S | Without Sterilization Case Double Wrap Only | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
| De Mayo V2 E Extension Arm 114 - EXT | Without Sterilization Case Double Wrap Only | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
|  |  |  |  |  |
| **Accessories** |  |  |  |  |
| 1 Clamp: 302, 903, 713-302, 713-717 or 114-CL1 Boot: 803-ABD or 803-CBD1 Distractor: 907 | 706-XSOR706-M | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |
| 1 Clamp: 302, 903, 713-302, 713-717 or 114-CL1 Boot: 803-ABD or 803-CBD1 Distractor: 907 | No Case Wrap Only | 270°F to 272°F132°C to 134° C | 4 minutes | 20 minutes |

**CAUTION: POSITIONERS MUST BE COOL BEFORE APPLYING TO THE PATIENT.**

Sterilization Parameters Certified by:

* Micro Test Laboratories (now Accuratus Lab Services)
* Accuratus Lab Services
* HighPower Validation Testing and Lab Services